## PTR
Post Test Review (PTR)
Unit test : | EM EUI SFT 2nd Run
IRR Minute Reference : | SOL-MN-ADSS-1000356487
IRR Title : | SOLAR ORBITER U-OTB AND ETB SWA ELECTRICAL INSTALLATION AND DRY RUN IRR
Previous PTR Minute Reference :
Previous PTR Title :
Actions from Previous Meeting (IRR or PTR)
Reference | Action Description | Actionee | Deadline | Disposition | Closure Status | Closure Date
IRR SOL-MN-ADSS-1000356487 Action 1 | Add step to check redundant relay box confing on ACS-296 | J Silverthorn | 2018-04-30 |  | Closed | 2018-04-30
IRR SOL-MN-ADSS-1000356487 Action 2 | There are a number of open issues with U-OTB. These shall be identify prior to test. If any issues shall be recorded as observations. | J Silverthorn | 2018-04-30 |  | Closed | 2018-04-30
General Results Chronology
Test Description | ACS Reference | As-Run Reference | From Date | To Date
EM SWA SFT HIS Integration and SFT Dry Run | ACS.SOL.OTB.196 |  | 2018-05-01 | 2016-06-08
Test Synthesis
Procedure Section | Section Title | OBS Raised during section (OBS number only) | Comments
ACS.SOL.OTB.196 Steps 6-16 | Off Load checks to check relay box operation
ACS.SOL.OTB.196 Steps 17 - 38 | SWA HIS in-rush measurements side A and side B | #1, #5
ACS.SOL.OTB.196 Steps 17-38 | SWA HIS sequence development | #3, #9
ACS.SOL.OTB.196 Steps 39, 40 | SWA PAS sequence delevopment | #6, #7
ACS.SOL.OTB.196 Steps 41,44 | SWA EAS sequence development | # 2, #4, #8
ACS.SOL.OTB.196 Step 52 | SWA SFT Side A | #8 | NCR to be raised
ACS.SOL.OTB.196 Step 54 | SWA SFT Side B | #8 | NCR to be raised
Observations
See Observation Sheet
Procedure Evolution
See Deviation Sheet
List of Open Tasks
Reference | Description of Task | Observation Reference | Status | Open Item Reference if applicable
Evolution Since IRR
| IRR Version | Current (PIR) Version | Justification
OC | V7 P21, sol patch 1, UG R1.0 | V7 P21, sol patch 1, UG R1.0 | N/A
SRDB | SISSOAI2_v5.540 inc IDB v5.1 | SISSOAI2_v3.310 | N/A
CSW | V2.1.3 Patch 4 | V2.1.3 Patch 4 | N/A
Results Directory
..\EUI SFT 2nd run
AOB
Note all values are currently checked using raw values and there will need to be an update when calibration curves are introduced. This should be developed from a database deliviery from EUI
Test Conclusion

## Observation Sheet
ACS Ref. | Observation Number | Step number(s) | Description | Disposition | Status | Questions/Clarifications
ACS.SOL.OTB.296 | 1 | Step 22 HIS switch On | at 14:06:30 : HIS_SPW_RX_FULL (5,3) events received. At 14:07:36 SWA_E_HIS_SAFE_ME (5,4) received | Recovery - switch HIS off. Attempt to enable heartbeat directly after switch on did not solve the problem. Grounding measurement between HISJ01 in 6 and the bench gives 366kOhm. With relays On in relay box (no LCL) this reduces to 317 kOhms Contact with ADS design support agrees to local grounding of pin 6 to the chassis plate. | Closed. | Temporary solution for OTB only. Will not occur on either ETB nor Spacecraft.
ACS.SOL.OTB.296 | 2 |  | TISS_EAS_ON failed to capture packet AB.TM.YIA58201 (also applicable to YIA58202) | Sequence updated . Incorrect tool was used to capture this HK packet. | Closed. | Tested in v1.2
ACS.SOL.OTB.296 | 3 |  | TISS_HIS_ON failed to capture packet YIA58213 | TM rate increased to 1sec / packet | Closed. | Action SWA team to update test procedure
ACS.SOL.OTB.296 | 4 |  | TISS_EAS_OFF TC ZIA58761 SWA_TC_EAS1_PWR_OFF rejected with 1,8 | Happened because EAS_ON was not fully completed (see OBS 2). Send TC to run transition macro AB.TC.ZIA58758. EAS could now be switched OFF | Closed.
ACS.SOL.OTB.296 | 5 |  | AB.TM.NIA01329 expected <=25.0 got 25.208 NIA01365 also failed <=25.0 got 25.037 Remove halt before switch On TC | Seqeunce TISS_HIS_ON modified for EM only to force all out-of limits to pass. Complete list of affected parameters passed to SWA by email. | Closed. | Action - SWA team to anayse and advise if any of these OOL measurements are abverse to EM operations.
ACS.SOL.OTB.296 | 6 |  | PAS_ON (5,3) SWA_E_SPU_SW_WARNING_RAISE YIA58957 received | SWA state that these are expecetd . | Closed. | Action ADS to update procedure for all expected TM (5,3) events.
ACS.SOL.OTB.296 | 7 |  | PAS_OFF TC ZIA58857 timeout is too short. Also TC ZIA58859 | Sequence updated to capture 1,7 event later. | Closed. | Tested in version 1.4 of TISS_PAS_OFF
ACS.SOL.OTB.296 | 8 |  | EAS_ON sequence : checks for NIA01503 can't be checked for now. (For both EAS1 and EAS2 TM packets YIA58904 and YIA58905 | NCR to be raised to track issue | Closed.

## Deviation Sheet
Step number(s) | Deviation | Justification
TIET_SWA_SFT v1.19 | Initial version
TIET_SWA_SFT v1.20 | update HV safety plug argument and abort if flight configuration is chosen | debug
TIET_SWA_SFT v1.21 | EM branch removed + proto associated | debug
TIET_SWA_SFT v1.22 | Remove dialog and correct the airsafe part for PAS | debug
TIET_SWA_SFT v1.23 | correct typo in report flow | debug
TIET_SWA_SFT v1.24 | add calltree | debug
TISS_HIS_ON v1.1 | Initial version
TISS_HIS_ON v1.4 | Implement A and B side option | debug
TISS_HIS_ON v1.5 | Remove halt to set heartbeat after HIS power up | debug
TISS_HIS_ON v1.6 | extend timeout for catching packet YIA58213 | OBS 3
TISS_HIS_ON v1.7 | Remove capture of first packet YIA58213 | OBS 3
TISS_HIS_ON v1.8-1.11 | Debugging of TM limits for EM SWA | OBS 5,9
TISS_HIS_ON v1.12 | Debugging of TM limits for EM SWA | OBS 5,9
TISS_HIS_ON v1.13 | Debugging of TM limits for EM SWA + remove halt before TC HIS on | OBS 5,9
TISS_HIS_OFF v1.1 | Initial version
TISS_PAS_ON v1.4 | Initial version
TISS_PAS_ON v1.5 | Update proto to include HV safety and forbid switch on if flight conf is chosen | debug
TISS_PAS_ON v1.6 | Update capture of packet for macro | debug
TISS_PAS_OFF v1.1 | Initial version
TISS_PAS_OFF v1.2 | Correct typo in title | debug
TISS_PAS_OFF v1.3 | update HV safety plug argument | debug
TISS_PAS_OFF v1.4 | add catch of 1,7 for switch off TC | OBS 7
TISS_EAS_ON v1.1 | Initial version
TISS_EAS_ON v1.2 | ACS-296, OBS 2: Update capture of packet YIA58201 and YIA58202 to HK capture | OBS 2
TISS_EAS_ON v1.3 | Correct typo in header
TISS_EAS_ON v1.4 | Correct typo in TM packets | OBS 8
TISS_EAS_ON v1.5 | Correct typo in TM packets | OBS 8
TISS_EAS_ON v1.6 | Correct typo in catch packets for EAS2 | OBS 8
TISS_EAS_ON v1.7-8 | Correct typo in catch packets for EAS1 and 2 | OBS 8
TISS_EAS_OFF v1.1 | Initial version
TISS_EAS_OFF v1.1 | Correct typo in titles/header | debug